Stopped Flow Fluorescence on APP

To Start APP:

- Turn on nitrogen tank (pressure ~120 psi-5 bar…should already be set)

- Check flow at end of tube at lamp housing

  - Wait at least 10 minutes to make sure housing is purged with nitrogen

- Turn lamp on (switch on back) and ignite (green button in back)

  - Check light at lamp housing (screw looking thing)

  - It takes 60 minutes for the lamp to warm up

- Moving right to left turn on machine

  - Channel 2 monochomater (in back)

    - Wait for monochromater to finish calibration routine-feel on top

  - Channel 1 monochomater (in back)

    - Wait for monochromater to finish calibration routine-feel on top

  - Sample handling unit (SHU) (in back)

  - Each of the three things under the computer (in back)

  - Acorn computer (in front)

  - NEXT…check screw on top of SHU (purple handle hex screwdriver)

    - Turn on water bath (2 switches)

    - Make sure door is open

To Run Vesicles:

***Remember that everything that is run through the machine will be diluted 1:1!!!  Make all solutions twice as concentrated to account for this***

- Open Xscan

  - Middle click—aquire

- Fill syringes with external buffer

  - Switches facing out—fill syringes below (in water bath)

  - Switches facing toward you—run to detector

- Click empty

  - Manually push up brown thing on left of SHU

  - Manually push up silver thing to refill brown thing

- Move switches to pointing toward the walls—refill inside syringes

- Repeat at least 4 times with external buffer

- Open SX18MV

  - Select new data (upper left corner of menu)

  - Chick on channel #1 button and change to channel #2

  - Click square button next to detector—will switch to emission

  - Set excitation wavelength to 452 nm for pyranine

    - Middle click on up arrow next to wavelength option

  - Click on button next to internal trigger and select external trigger

  - Click up arrow next to range until it is at –0.10 – 0.10

  - On graph, click near zero to move line so “volts in” is as close to zero as possible (should get to +/- 0.001)

  - Click down arrow to set range back to –5 to 5

  - Entrance and exit slit width is 2mm—can be adjusted to give a better signal

- Empty right syringe and fill with vesicles (pump up and down)

- Run a small amount past the detector—click empty and manually push the silver thing up

- The reading should zoom up—(a) change range and (b) click graph so line is ~3/4 from top of graph

- Click acquire a couple of times to run vesicles past detector

- Looking at data

  - Click display

  - Click on a run

  - Click on overlay to view multiple runs at once

  - Click zoom—left click on graph above run on left side, make a box and then right click

  - If runs are on top of one another, proceed.  If not, continue acquiring 0.01 sec runs until they are.

- Collect baseline

  - Click arrow next to time and set to desired time

  - Click acquire

    - If fluorescence decreases, machine is not clean.  Empty syringes and flush with lots of water then ethanol then water then buffer

    - NEVER run ethanol right after running sucrose buffer!!!

  - (a) Click shots and change to 3, (b) click average to turn on and (c) click oversampling to turn on, (d) click aquire.

    - This will take lots of data points and the average of three runs for a good baseline.

- Make group (a save folder)

  - Click disc

  - Click make group

- Rename run

  - Click on disc

  - Click on the run you want to rename

  - Click on rename—be sure to write the name of the group and run in your lab notebook!

- Save run

  - Chick on a run on the right side

  - Click on disc

  - Click on savefile

  - IT IS BEST TO SAVE EVERYTHING RIGHT AWAY BECAUSE THE ACORN LIKES TO LOCK UP!!!

- Looking at data

  - Click display

  - Click on a run

  - Click on overlay (to turn it on) to view multiple runs at once

  - Click zoom—left click on graph above run on left side, make a box and then right click

  - clearscr will clear all the data off the screen

- Load left syringe with drug, etc.

- Run past detector

  - Click shots and change back to 1

  - At least 3 short runs 0.1 secs

  - Overlay data under display to make sure they are the same

  - Chose desired time (same as baseline)

  - Click shots and change to 3

  - Click aquire

- Rinse left syringe with buffer between runs

  - Have switch for vesicles facing toward wall and switch for left syringe facing toward you

- Clean machine when finished

  - Exit program

  - Turn off lamp—next to wall

  - Open Xscan

  - Repeat steps 1-4

    - At least 2 syringes of water

    - At least 2 syringes of ethanol

    - At least 2 syringes of water

    - Shortcutting this leads to more work later and you’ll regret it

- Shut off everything starting with the water bath and working backwards

- Leave nitrogen on for about 10 minutes to flush lamp housing.

- Loosen screw on SHU to drain water from machine.

- Turn nitrogen off